## Volume
| alumina | Al2O3 | 3.95 |  |  |  |  | 0.63 | JLOX-101 data sheet
| silica gel | SiO2 | 0.7 |  |  |  |  | 0.64 | http://www.adsorbentcn.com/13x.htm
| zeolite | AlSiO | 0.63 |  | Li-X, Na-X | 0.63 | kg/m^3 | 0.79 | https://www.ncbi.nlm.nih.gov/pmc/articles/PMC5575677/
| dia-in | len-in | id | length | liters | columns | kg | kgtotal | lbs |  | gas volume
[13 empty] | 2.5 | 10
| 1.37 | 12 | 3.4798 | 30.48 | 0.28988 | 1 | 0.18262 | 0.18262 | 0.40177 | AlO3 Tester | 0.086963 |  | 1.2823 | 12.8232
| 1.37 | 18 | 3.4798 | 45.72 | 0.43482 | 1 | 0.27393 | 0.27393 | 0.60265 |  | 0.13044
| 1.37 | 25 | 3.4798 | 63.5 | 0.60391 | 1 | 0.38046 | 0.38046 | 0.83702 | Alt Tester 5 lpm | 0.18117
| 1.37 | 37 | 3.4798 | 93.98 | 0.89379 | 1 | 0.56309 | 0.56309 | 1.2388 |  | 0.26814
| 1.872 | 14 | 4.7549 | 35.56 | 0.63144 | 1 | 0.39781 | 0.39781 | 0.87517 |  | 0.18943
| 1.872 | 20 | 4.7549 | 50.8 | 0.90205 | 1 | 0.56829 | 0.56829 | 1.2502 | Tester Column 10 lpm | 0.27062
| 1.872 | 26 | 4.7549 | 66.04 | 1.1727 | 1 | 0.73878 | 0.73878 | 1.6253 |  | 0.3518 | 34.798 | 1440 | min/day
| 2.87 | 20 | 7.2898 | 50.8 | 2.1202 | 2 | 1.3358 | 2.6715 | 5.8773 | O24U Dryer Column | 0.63607 |  | 1000 | l/m3
| 2.87 | 24 | 7.2898 | 60.96 | 2.5443 | 2 | 1.6029 | 3.2058 | 7.0528 | O24U PSA Column | 0.76329
| 2.87 | 38 | 7.2898 | 96.52 | 4.0285 | 2 | 2.5379 | 5.0759 | 11.1669 |  | 1.2085 |  | 20 | sm3/day
[13 empty] | 13.8889 | l/min
| 3.37 | 10 | 8.5598 | 25.4 | 1.4617 | 1 | 0.92086 | 0.92086 | 2.0259 |  | 0.4385
| 3.37 | 11 | 8.5598 | 27.94 | 1.6078 | 1 | 1.0129 | 1.0129 | 2.2285 | Short Tank | 0.48235
| 3.37 | 20 | 8.5598 | 50.8 | 2.9234 | 1 | 1.8417 | 1.8417 | 4.0518 |  | 0.87701
| 3.37 | 21 | 8.5598 | 53.34 | 3.0695 | 1 | 1.9338 | 1.9338 | 4.2544 | Long Tank | 0.92086
[12 empty] | volume | 4.5 | l
| 3.87 | 12 | 9.8298 | 30.48 | 2.3131 | 1 | 1.4573 | 1.4573 | 3.206 |  |  | pmax | 50 | kPa | 0.5 bar | 7.5 psig
| 3.87 | 18 | 9.8298 | 45.72 | 3.4696 | 1 | 2.1859 | 2.1859 | 4.8089 |  | 1.0409 | pmin | 30 | kPa | 0.3 bar
| 3.87 | 36.5 | 9.8298 | 92.71 | 7.0357 | 1 | 4.4325 | 4.4325 | 9.7514 |  | 2.1107 | Stored Vol | 0.9 | l
| 3.87 | 48 | 9.8298 | 121.92 | 9.2524 | 1 | 5.829 | 5.829 | 12.8238 |  | 2.7757 | Delivery | 5 | l/min
[12 empty] | Time | 10.8 | sec
| 3.068 | 17 | 7.7927 | 43.18 | 2.0594 | 2 | 1.2975 | 2.5949 | 5.7088 |  | 0.61783
[12 empty] | packing factor | 0.7
| 2.87 | 12 | 7.2898 | 30.48 | 1.2721 | 2 | 0.80145 | 1.6029 | 3.5264 | Slice Proto
| 5.87 | 12 | 14.9098 | 30.48 | 5.3217 | 1 | 3.3527 | 3.3527 | 7.3758 |  | 4.856
| 5.87 | 36.5 | 14.9098 | 92.71 | 16.1868 | 1 | 10.1977 | 10.1977 | 22.4349 | Big Slice 100 lpm @2
| 7.87 | 24 | 19.9898 | 60.96 | 19.1316 | 1 | 12.0529 | 12.0529 | 26.5164 | 100 lpm each
[13 empty] | "200 lpm, 1100W, USD270, 20kg" pump |  |  |  | 550W 2.8A
| 2.37 | 20 | 6.0198 | 50.8 | 1.4458 | 2 | 0.91087 | 1.8217 | 4.0078 | MVP slice 20-30 lpm @2 |  | in | 0.09 | 0.09 | 0.09 | 0.09 | 0.05 | top | 0.09
[13 empty] | 0.965 | 0.965 | 0.965 | 0.965 | 0.505 | bot | 0.965
[13 empty] | 2.5 | 2.5 | 2.5 | 2.5 | 2.5 | dia | 3
| 4.87 | 36.5 | 12.3698 | 92.71 | 11.1415 | 1 | 7.0191 | 7.0191 | 15.4421
[12 empty] | cm | 0.2286 | 0.2286 | 0.2286 | 0.2286 | 0.127 | top | 0.2286
[13 empty] | 2.4511 | 2.4511 | 2.4511 | 2.4511 | 1.2827 | bot | 2.4511
|  |  | 7.1 | 44 | 1.742 | 2 | 1.0975 | 2.195 | 4.8289 | Oxikit Li sieve bed |  |  | 6.35 | 6.35 | 6.35 | 6.35 | 6.35 | dia | 7.62
|  |  | 7.1 | 52 | 2.0588 | 2 | 1.297 | 2.5941 | 5.7069 | Oxikit Li o2 tank
[12 empty] | stroke | 2.2225 | 2.2225 | 2.2225 | 2.2225 | 1.1557 | cm | 2.2225
[12 empty] | area | 31.6692 | 31.6692 | 31.6692 | 31.6692 | 31.6692 | cm^2 | 45.6037
| 1.37 | 12 | 3.4798 | 30.48 | 0.28988 | 1 | 0.18262 | 0.18262 | 0.40177 | puffer 5-7 lpm pulse |  | displacement | 70.3848 | 70.3848 | 70.3848 | 70.3848 | 36.6001 | cm^3 | 101.3542
[12 empty] | Cylinders | 2 | 2 | 2 | 4 | 2 |  | 4
| 1.38 | 18 | 3.5052 | 45.72 | 0.44119 | 1 | 0.27795 | 0.27795 | 0.61148 | 1.25 sch 40
| 1.61 | 18 | 4.0894 | 45.72 | 0.6005 | 1 | 0.37832 | 0.37832 | 0.8323 | 1.5 sch 40 |  | RPM | 1400 | 1680 | 2000 | 1100 | 1680 | 60 hz | 1700
| 2.067 | 10 | 5.2502 | 25.4 | 0.54989 | 1 | 0.34643 | 0.34643 | 0.76214 | 2 sch 40
| 2.067 | 6 | 5.2502 | 15.24 | 0.32993 | 1 | 0.20786 | 0.20786 | 0.45728 |  |  | FA Volume | 197.0775 | 236.493 | 281.5393 | 309.6933 | 122.9764 | lpm | 689.2083
[12 empty] | Static CR | 10.7222 | 10.7222 | 10.7222 | 10.7222 | 10.1 |  | 10.7222
[13 empty] | 7.0385 | 8.4462 | 10.055 | 11.0605 | 4.392 | CFM FA | 24.6146
| TesterData |  |  | 180 Bends | 2 | 2.5 | 3 | 4 | 6 |  |  |  | 5.6308 | 6.7569 | 8.044 | 8.8484 | 3.5136 | CFM net | 19.6917
| kg sieve | 1.1 |  | Radius | 50 | 78.75 | 85 | 113.05 | 169.575
| lpm net | 10 |  | Bend length | 157.0796 | 247.4004 | 267.0354 | 355.157 | 532.7356 |  |  | VPSA single | VPSA N+1 | PSA + Purge
| O2 recovery | 0.6 |  | Tail length | 10 | 150 | 20 | 21 | 22 |  | Recovery | 0.9 | 0.9 | 0.6
| Zeolite utilization | 15.1515 | l/(min*kg) | Total path | 177.0796 | 547.4004 | 307.0354 | 397.157 | 576.7356 |  | Cycle Factor | 0.67 | 1 | 1
|  |  |  | (in) | 6.9716 | 14.5 | 12.088 | 15.6361 | 22.7061
[11 empty] | DC | ZW280 | 2680N50 | 2680 2:1 | Welch 2585 | 200 lpm | 10CFM | 12cfm
[6 empty] | Recovery | 0.9 | Forward O2 recovery |  | Air lpm | 30 | 75 | 120.96 | 171 | 160 | 192 | 201.6 | 235.2 | l/min
[6 empty] | Cycle Factor | 0.5 | total-regen/total |  | O2 lpm | 2.835 | 7.0875 | 11.4307 | 16.1595 | 15.12 | 18.144 | 19.0512 | 22.2264 | l/min
| alumina | Al2O3 | 3.95 | g/cc |  | lpm/kg raw | 15.1515 | Zeolite utilization |  | Zeolite sieve | 0.37422 | 0.93555 | 1.5089 | 2.1331 | 1.9958 | 2.395 | 2.5148 | 2.9339 | kg
| silica gel | SiO2 | 0.7 |  |  | fill factor | 0.8 | sieve/volume |  | Zeolite max | 0.46777 | 1.1694 | 1.8861 | 2.6663 | 2.4948 | 2.9938 | 3.1434 | 3.6674 | kg
| zeolite | AlSiO | 0.63 |  |  | Li-X, Na-X | 0.63 | g/cc |  | litres | 0.7425 | 1.8563 | 2.9938 | 4.2322 | 3.96 | 4.752 | 4.9896 | 5.8212 | l
| dia-in | len-in | id | length | liters | columns | kgmax | kg-sieve | lbs |  | Throttle | O2 | Air | CFM
| 2.067 | 20 | 5.2502 | 50.8 | 1.0998 | 1 | 0.69286 | 0.55428 | 1.2194 |  | 1 | 4.1991 | 44.4352 | 1.587 | DC unit/ tiny AC
| 1.87 | 31 | 4.7498 | 78.74 | 1.3952 | 1 | 0.87897 | 0.70318 | 1.547 |  | 1 | 5.3271 | 56.3716 | 2.0133 | 3.3 cfm ZW280
| 2.37 | 38.5 | 6.0198 | 97.79 | 2.7832 | 1 | 1.7534 | 1.4027 | 3.086 |  | 1 | 10.6269 | 112.4536 | 4.0162 | Thomas 2660/E160
| 2.87 | 48 | 7.2898 | 121.92 | 5.0886 | 1 | 3.2058 | 2.5646 | 5.6422 | 20+ lpm Li-X | 1 | 19.4291 | 205.5989 | 7.3428 | Demo U-tube 3/4 HP + AMR
| 3.87 | 36 | 9.8298 | 91.44 | 6.9393 | 1 | 4.3718 | 3.4974 | 7.6943 | 25+ lpm Li-X | 1 | 26.4955 | 280.3756 | 10.0134
| 5.87 | 24 | 14.9098 | 60.96 | 10.6434 | 1 | 6.7053 | 5.3643 | 11.8014 | 25+ lpm mixed media | 1 | 40.6383 | 430.0345 | 15.3584 | Demo 20-30 lpm @550W
| 5.87 | 36 | 14.9098 | 91.44 | 15.965 | 1 | 10.058 | 8.0464 | 17.702 | 50 lpm Li-X | 1 | 60.9574 | 645.0517 | 23.0376 | Demo 50-60 lpm @1100W |  |  | 806.3146 | 3450
| 7.87 | 36 | 19.9898 | 91.44 | 28.6974 | 1 | 18.0794 | 14.4635 | 31.8197 | 100 lpm Li-X | 1 | 109.572 | 1159.4921 | 41.4104 |  |  | disp | rpm
| tanks | [9 empty] | Target system at scale - 50 l/m per slice, 2 slice minimum | [5 empty] | 0.5 | 1612.6293 | 2.1394
| 3.37 | 21 | 8.5598 | 53.34 | 3.0695 | Long Tank |  |  |  |  | 6" column with mixed sieve media | [5 empty] | 0.3 | 2687.7155 | 1.2836
| 3.87 | 24 | 9.8298 | 60.96 | 4.6262 | 1 | 2.9145 | 2.3316 | 5.1295 |  | 12 CFM pump + smaller vacuum pump for 1:1 timing
| 5.87 | 18 | 14.9098 | 45.72 | 7.9825 | [5 empty] | N+1 with N>=2
| 5.87 | 48 | 14.9098 | 121.92 | 21.2867
| 7.87 | 48 | 19.9898 | 121.92 | 38.2632 | [5 empty] | 2/3 size pump per slice, shared vacuum use, smaller slices, time-slice
3 slice minimum?
|  | Press | [8 empty] | 2X vacuum rate for 1/2 time?
Pump | Press | Press |  | Vac | [6 empty] | 1.5X with both pumps
Column | Prod | 2X Prod | BD | Vac
[11 empty] | 160 + 80 | 120 + 120
| single-pump slice | [9 empty] | 3+1, 2 | 3 time slices
Pa | [9 empty] | 1x
Pb | [9 empty] | 1x
[12 empty] | (120+120)*1/3 + 120*2/3 |  | 160
Ca | 1 | 1 | 1 |  |  |  | 1 | 1 | 1 | 3/6 |  | 1/3*240 + 2/3*120
Cb |  |  |  | 1 | 1 | 1 |  |  |  | 3/6 |  |  | 120*2/3 | 80
| 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1x
Pa | [9 empty] | 6/5x |  | 2X 2660/143(114)
Pb | [9 empty] | 6/5x |  | 171 | Axg 2:1
32.3 lpm O2
Ca | 1 |  |  | 1 | 2 | 1 | 1 |  |  | 5/6
Cb | 1 | 2 | 1 | 1 |  |  | 1 | 2 | 1 | 5/6 |  | 228
| 2 | 2 | 1 | 2 | 2 | 1 | 2 | 2 | 1 | 5/3 |  | 57 | lpm offset
[12 empty] | 5 | s period
Pa | [9 empty] | 4/3x |  | 2.375 | l carry over
Pb | [9 empty] | 4/3x
5 liter, +50kPa
Ca |  |  | 2 | 1 |  |  | 2 | 1 |  | 3/4
Cb | 2 | 1 |  |  | 2 | 1 |  |  | 2 | 3/4
Prod | 2 | 1 | 2 | 1 | 2 | 1 | 2 | 1 | 2 | 3/2
Pa | [9 empty] | 3/2x
Pb | [9 empty] | 3/2x
Ca |  |  |  | 2 | 1 | 1 |  |  |  | 2/3
Cb | 2 | 1 | 1 |  |  |  | 2 | 1 | 1 | 2/3
| 2 | 1 | 1 | 2 | 1 | 1 | 2 | 1 | 1 | 4/3
dual-pump slice
Pa | [9 empty] | 1x
Pb | [9 empty] | 1x
Va | [9 empty] | 1/2x
Vb | [9 empty] | 1/2x
Ca | 2 | 2 | 2 |  | 1 | 1 | 2 | 2 | 2 | 6/6
Cb |  | 1 | 1 | 2 | 2 | 2 |  | 1 | 1 | 6/6
| 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 12/6
Pa | [9 empty] | 1x - 4/3x
Pb | [9 empty] | 1x
Pc | [9 empty] | 1x
Va | [9 empty] | 1/2x
Vb | [9 empty] | 1/2x
Vc | [9 empty] | 1/2x
Ca | 1.5 | 3 | 1.5 |  | 1.5 | 1.5 | 1.5 | 3 | 1.5 | 6/6
Cb | 1.5 | 1.5 | 1.5 | 3 | 1.5 |  | 1.5 | 1.5 | 1.5 | 6/6
Cc | 1.5 |  | 1.5 | 1.5 | 1.5 | 3 | 1.5 |  | 1.5 | 6/6
| 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3
Ca | 2 | 2 | 2 |  | 1.5 | 1.5 | 2 | 2 | 2 | 6/6
Cb | 1.5 | 1.5 | 2 | 2 | 2 |  | 1.5 | 1.5 | 2 | 6/6
Cc | 2 |  | 1.5 | 1.5 | 2 | 2 | 2 |  | 1.5 | 6/6
| 4 | 2 | 4 | 2 | 4 | 2 | 4 | 2 | 4
| 1.75 | 2.5 | 7 | 10 | 7 | 10
| 3.5 | 2.5 | 14 | 10 | 14 | 10
| 1.67 | 2.67 | 5 | 8 | 5 | 8
| 3.33 | 2.67 | 10 | 8 | 10 | 8
Pa | [9 empty] | 1x
Pb | [9 empty] | 1x
Va | [9 empty] | 1/2x
Vb | [9 empty] | 1/2x
Ca | 1 | 2 | 2 | 1 | 1 | 1 | 2 | 2 | 1 | 5/5
Cb | 1 | 1 | 1 | 2 | 2 | 1 | 1 | 1 | 2 | 5/5
| 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 10/5
Pa | [9 empty] | 1x
Pb | [9 empty] | 1x
Va | [9 empty] | 1/2x
Vb | [9 empty] | 1/2x
Ca | 2 | 1 | 2 | 2 | 1 | 2 | 2 | 1 | 2 | 5/5
Cb | 2 | 1 | 2 | 2 | 1 | 2 | 2 | 1 | 2 | 5/5
| 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 2 | 10/5

## Compressor
HP = [144 N P1 V k / (33000 (k - 1))] [(P2 / P1)(k - 1)/N k - 1]
where
HP = horsepower
N = number of compression stages
k = 1.41 = adiabatic expansion coefficient
P1 = absolute initial atmospheric pressure (psi) (14.7 psi at sea level)
P2 = absolute final pressure after compression (psi)
V = volume of air at atmospheric pressure - standard or free air (scfm, ft3/min)
| N | 1 | Stages |  | N | 1 | 1 | Stages
| k | 1.41 | constant |  | k | 1.41 | 1.41 | constant
| P1 | 14.7 | PSIA |  | P1 | 14.7 | 1.5 | PSIA
| P2 | 60 | PSIA |  | P2 | 30 | 14.7 | PSIA
| V | 5 | SCFM |  | V | 12 | 7 | SCFM
| Pg | 45.3 | PSIG |  | Pg | 45.3 | 45.3 | PSIG
| HPideal | 0.55732 |  |  | HPideal | 0.6102 | 0.14842
| eff | 0.8 |  |  | eff | 0.8 | 0.8
| KW | 0.51949 |  |  | KW | 0.56878 | 0.13834
| N | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| k | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41
| P1 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7
| P2 | 29.7 | 29.7 | 29.7 | 29.7 | 29.7 | 29.7 | 29.7 | 29.7 | 29.7 | 29.7
| V | 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16 | 18 | 20
| Pg | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15 | 15
| HPideal | 0.10012 | 0.20023 | 0.30035 | 0.40046 | 0.50058 | 0.60069 | 0.70081 | 0.80092 | 0.90104 | 1.0012
| eff | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| KW | 0.09332 | 0.18664 | 0.27996 | 0.37328 | 0.4666 | 0.55992 | 0.65324 | 0.74656 | 0.83988 | 0.9332
| N | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| k | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41
| P1 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7
| P2 | 37.2 | 37.2 | 37.2 | 37.2 | 37.2 | 37.2 | 37.2 | 37.2 | 37.2 | 37.2
| V | 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16 | 18 | 20
| Pg | 22.5 | 22.5 | 22.5 | 22.5 | 22.5 | 22.5 | 22.5 | 22.5 | 22.5 | 22.5
| HPideal | 0.13674 | 0.27348 | 0.41023 | 0.54697 | 0.68371 | 0.82045 | 0.9572 | 1.0939 | 1.2307 | 1.3674
| eff | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| KW | 0.12746 | 0.25492 | 0.38238 | 0.50984 | 0.6373 | 0.76477 | 0.89223 | 1.0197 | 1.1471 | 1.2746
| N | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1
| k | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41 | 1.41
| P1 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7 | 14.7
| P2 | 44.7 | 44.7 | 44.7 | 44.7 | 44.7 | 44.7 | 44.7 | 44.7 | 44.7 | 44.7
| V | 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16 | 18 | 20
| Pg | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30 | 30
| HPideal | 0.16845 | 0.33689 | 0.50534 | 0.67379 | 0.84223 | 1.0107 | 1.1791 | 1.3476 | 1.516 | 1.6845
| eff | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8 | 0.8
| KW | 0.15701 | 0.31403 | 0.47104 | 0.62805 | 0.78507 | 0.94208 | 1.0991 | 1.2561 | 1.4131 | 1.5701
|  | 2 | 4 | 6 | 8 | 10 | 12 | 14 | 16 | 18 | 20
| 15 psig/vac | 0.09332 | 0.18664 | 0.27996 | 0.37328 | 0.4666 | 0.55992 | 0.65324 | 0.74656 | 0.83988 | 0.9332
| 25 psig/vpsa | 0.12746 | 0.25492 | 0.38238 | 0.50984 | 0.6373 | 0.76477 | 0.89223 | 1.0197 | 1.1471 | 1.2746
| 40 psig/psa | 0.15701 | 0.31403 | 0.47104 | 0.62805 | 0.78507 | 0.94208 | 1.0991 | 1.2561 | 1.4131 | 1.5701
Prototype pdc scenarios
| rpm | st. disp | disp | pump eff | slpm | scfm | psig | hp | Nstages | k | P1 | P2 | V | Pg | HP | eff | kw
| 1750 | 0.3 | 525 | 0.8 | 420 | 14.841 | 22 | 0.99785 | 1 | 1.41 | 14.7 | 36.7 | 14.841 | 22 | 0.99785 | 0.8 | 0.93012
| 1200 | 0.3 | 360 | 0.8 | 288 | 10.1767 | 22 | 0.68424 | 1 | 1.41 | 14.7 | 36.7 | 10.1767 | 22 | 0.68424 | 0.8 | 0.6378
| 2880 | 0.5 | 1440 | 0.8 | 1152 | 40.7067 | 22 | 2.737 | 1 | 1.41 | 14.7 | 36.7 | 40.7067 | 22 | 2.737 | 1.8 | 1.1339
| 3450 | 0.5 | 1725 | 0.8 | 1380 | 48.7633 | 20 | 3.0518 | 1 | 1.41 | 14.7 | 34.7 | 48.7633 | 20 | 3.0518 | 1.8 | 1.2643
| 120 | 16: 120
| scfh | 2: psig; 4: 120 lpm; 16: scfh; 18: in.; 20: 120 lpm
| 250 | 0 | 117.9167 | 4.1667 |  |  |  |  | 220 | 3 | 103.7667
| 200 | 7.5 | 94.3333 | 3.3333 |  |  |  |  | 200 | 6 | 94.3333
| 150 | 15 | 70.75 | 2.5 |  |  |  |  | 150 | 13 | 70.75
| 120 | 22.5 | 56.6 | 2 |  |  |  |  | 100 | 19 | 47.1667
[9 empty] | 50 | 23 | 23.5833
| zw280 | 16: 5; 18: 26.5; 20: 2.3583
| scfh | 2: psig; 4: zw280; 16: zw280
| 220 | 2: 0; 4: 103.7667; 16: scfh; 18: in.; 20: zw280
| 170 | 2: 7.5; 4: 80.1833; 16: 200; 18: 3; 20: 94.3333
| 150 | 2: 15; 4: 70.75; 16: 170; 18: 6; 20: 80.1833
| 120 | 2: 22.5; 4: 56.6; 16: 150; 18: 9; 20: 70.75
| Ilmvac | 16: 100; 18: 14; 20: 47.1667
| scfh | 2: psig; 4: 2585b; 16: 50; 18: 19; 20: 23.5833
| 430 | 0 | 202.8167 | 7.1667 |  |  |  |  | 5 | 25 | 2.3583
| 390 | 2: 7.5; 4: 183.95; 6: 6.5
| 340 | 2: 15; 4: 160.3667; 6: 5.6667; 16: 2585b
| 310 | 22.5 | 146.2167 | 5.1667 |  |  |  |  | scfh | in. | 2585b
[9 empty] | 340 | 7 | 160.3667
[9 empty] | 300 | 9 | 141.5
[9 empty] | 250 | 13 | 117.9167
[9 empty] | 200 | 17 | 94.3333
[9 empty] | 150 | 20 | 70.75
[9 empty] | 100 | 22 | 47.1667
[9 empty] | 50 | 24 | 23.5833
[9 empty] | 5 | 25.5 | 2.3583

## Convection
Convection
https://www.engineeringtoolbox.com/convective-heat-transfer-d_430.html
Newton's equation for convective heat transfer
q = hc A dT
Airflow over a surface
hcW = 12.12 - 1.16 v + 11.6 v1/2
|  | Finned split tube exchanger | Radiator convoluted fin tube | 30 row exchanger | 40 row exchanger
| Diameter | 4.2 | 1.2 |  |  | cm
| Length | 38 | 400 | 27 | 27 | cm
| Section Area | 13.8544 | 1.131 | 30 | 40 | cm^2
| Volume | 526.4681 | 452.3893 | 810 | 1080 | cm^3
| Flow rate | 75 | 160 | 200 | 550 | slpm
| Pressure | 150 | 150 | 150 | 150 | kPa
| Flow rate | 1250 | 2666.6667 | 3333.3333 | 9166.6667 | sccm/sec
| Volumetric | 500 | 1066.6667 | 1333.3333 | 3666.6667 | accm/Sec
|  | 36.0896 | 943.1404 | 44.4444 | 91.6667 | cm/sec
| Velocity | 0.3609 | 9.4314 | 0.44444 | 0.91667 | m/sec
| hc | 18.67 | 36.8039 | 19.3378 | 22.1628 | W/m^2*C
| surface area | 501.3982 | 1507.9645 | 8500 | 11350 | cm^2
| area | 0.05014 | 0.1508 | 0.85 | 1.135 | m^2
| internal air | 45 | 45 | 45 | 45 | C
| surface | 25 | 25 | 25 | 25 | C
| q | 18.7222 | 110.9978 | 328.7422 | 503.096 | W

## Volume-radialflow
Radial inflow (centripetal) design from the paper |  |  | 3-D Modeling of Gas–Solid Two-Phase Flow in a Pi-Shaped Centripetal Radial Flow Adsorber
density | 0.63
Rinner | 1.5 | cm
Router | 7.8 | cm
height | 18.7 | cm
area | 184.0659 | cm^2
volume | 3442.0326 | ml
weight | 2168.4805 | g

## isotherm
O2
| 32 | g/mol
| 22.4 | l/mol
| 1.4286 | g/l
1.4291
| 100 | lpm
| 205.7904 | kg/day
| 485.9313 | lpm = 1MTpd
N2
| 28 | g/mol
| 22.4 | l/mol
| 1.25 | g/l
|  | full range | oxikit | hp | vpsa |  |  |  | Pressure | 1.5 | bar
| adsL | 0 | 0.65 | 1.1 | 0.1 | mmol/g
|  | 0 | 18.2 | 30.8 | 2.8 | mg/g
|  | 0 | 14.56 | 24.64 | 2.24 | ml/g
| adsH | 1.4 | 1.2 | 1.4 | 0.8 | mmol/g
|  | 39.2 | 33.6 | 39.2 | 22.4 | mg/g
|  | 31.36 | 26.88 | 31.36 | 17.92 | ml/g
|  | 31.36 | 12.32 | 6.72 | 15.68 | l/kg
| P25 | P100 |  | 6" 1Li:3Na | 4" Li-X | big p 1:3 | big pump | good ratio | 50/50 | 3" LiNaX 1:1 | good ratio | 50/50 | tester | tester
Bed weight | 2.5 | 15 | kg | 5 | 3.89 | 1.28 | 2.56 | 2.56 | 2.56 | 1.6 | 0.7 | 0.7 | 0.55 | 0.55 | kg
N2 adsorbed | 39.2 | 235.2 | l | 78.4 | 60.9952 | 20.0704 | 40.1408 | 40.1408 | 40.1408 | 25.088 | 10.976 | 10.976 | 8.624 | 8.624 | l
Air rate | 280 | 1152 | l/min | 550 | 550 | 264 | 264 | 210 | 192 | 192 | 85 | 75 | 100 | 50 | l/min
N2 Rate | 221.2 | 910.08 | l/min | 434.5 | 434.5 | 208.56 | 208.56 | 165.9 | 151.68 | 151.68 | 67.15 | 59.25 | 79 | 39.5 | l/min
N2 ads time | 0.17722 | 0.25844 | minutes | 0.18044 | 0.14038 | 0.096233 | 0.19247 | 0.24196 | 0.26464 | 0.1654 | 0.16345 | 0.18525 | 0.10916 | 0.21833 | minutes
| 10.6329 | 15.5063 | seconds | 10.8262 | 8.4228 | 5.774 | 11.548 | 14.5175 | 15.8785 | 9.9241 | 9.8073 | 11.1149 | 6.5499 | 13.0997 | seconds
Bed Util | 0.9 | 0.9 |  | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 | 0.9 |  |  | PSA
O2 volume | 9.3782 | 56.2694 | l | 18.7565 | 14.5925 | 4.8017 | 9.6033 | 9.6033 | 9.6033 | 6.0021 | 2.6259 | 2.6259 | 2.0632 | 2.0632 | l | Bed weight | 1.7 | kg
[16 empty] | N2 adsorbed | 20.944 | l | PSA
[16 empty] | Air rate | 160 | l/min
Desorb rate | 224 | 921.6 | i/min | 425 | 425 | 230 | 230 | 180 | 150 | 150 | 120 | 60 | 80 | 50 | i/min | N2 Rate | 126.4 | l/min
Desorb time | 0.175 | 0.25521 | minutes | 0.18447 | 0.14352 | 0.087263 | 0.17453 | 0.223 | 0.26761 | 0.16725 | 0.091467 | 0.18293 | 0.1078 | 0.17248 | minutes | N2 ads time | 0.1657 | minutes
| 10.5 | 15.3125 | seconds | 11.0682 | 8.6111 | 5.2358 | 10.4715 | 13.3803 | 16.0563 | 10.0352 | 5.488 | 10.976 | 6.468 | 10.3488 | seconds |  | 9.9418 | seconds
[16 empty] | O2 recovery | 0.6
Cycle time | 0.35222 | 0.51365 | minutes | 0.36491 | 0.2839 | 0.1835 | 0.36699 | 0.46496 | 0.53225 | 0.33265 | 0.25492 | 0.36818 | 0.21696 | 0.39081 | minutes | O2 volume | 3.3404 | l
| 21.1329 | 30.8188 | seconds | 21.8945 | 17.0339 | 11.0097 | 22.0195 | 27.8977 | 31.9348 | 19.9593 | 15.2953 | 22.0909 | 13.0179 | 23.4485 | seconds | Cycle time | 0.1657 | minutes
[17 empty] | 9.9418 | seconds
Product rate | 26.6264 | 109.5487 | l/min | 51.4005 | 51.4005 | 26.1676 | 26.1676 | 20.6539 | 18.043 | 18.043 | 10.3008 | 7.1321 | 9.5094 | 5.2793 | l/min
[16 empty] | Product rate | 20.16 | l/min
output duty | 50% | 50% |  | 50% | 50% | 50% | 50% | 50% | 50% | 50% | 50% | 50% | 50% | 50%
output time | 0.088608 | 0.12922 |  | 0.090219 | 0.07019 | 0.048117 | 0.096233 | 0.12098 | 0.13232 | 0.0827 | 0.081727 | 0.092624 | 0.054582 | 0.10916
carryover | 7.0189 | 42.1136 | liters | 14.1192 | 10.9847 | 3.5426 | 7.0851 | 7.1046 | 7.2158 | 4.5099 | 1.784 | 1.9653 | 1.5442 | 1.4869 | liters
Column Dia | 7.3025 | 14.9225 | mm | 14.9225 | 9.8425 | 7.3025 | 7.3025 | 7.3025 | 7.3025 | 7.3025 | 4.7625 | 4.7625 | 4.7625 | 4.7625 | mm
Column area | 41.8825 | 174.8933 | cm^2 | 174.8933 | 76.0853 | 41.8825 | 41.8825 | 41.8825 | 41.8825 | 41.8825 | 17.8139 | 17.8139 | 17.8139 | 17.8139 | cm^2
Vmt | 2.6741 | 2.6348 | kcm/min | 1.2579 | 2.8915 | 2.5213 | 2.5213 | 2.0056 | 1.8337 | 1.8337 | 1.9086 | 1.6841 | 2.2454 | 1.1227 | kcm/min
Vmt | 0.44569 | 0.43913 | m/sec | 0.20965 | 0.48192 | 0.42022 | 0.42022 | 0.33427 | 0.30562 | 0.30562 | 0.3181 | 0.28068 | 0.37424 | 0.18712 | m/sec

## Model-O24U
Bed-product recovery ratio
N2 | 0.78 |  |  | Waste | 0.79 | .78 N2 + .01 O2
O2 | 0.209 | 95% |  | Product | 0.21 | 95% O2, 5% Ar
Ar | 0.009 | 4.091%
Other | 0.002 | 0.9091% |  |  | 19.86% | recovered in product
| 1 | 0.22 |  |  | 1.145% | adsorbed in bed
|  | 4.5455 | air/O2
N2cap | 22 | l/kg | Li-X ~= 22l/kg
iso-1 | 0.72 |  | fraction of capacity used from P/P0 range read off iostherm chart for adsorb/desorb pressure ratio
cyclecap | 15.84 | l/kg | liters of waste per kg of adsorbent, adsorbed each cycle, no hysteresis
bedweight | 1.8 | kg
bedcap | 28.512 | l | liters of waste adsorbed in the bed each cycle
Prodrate | 20 | l/min
Prodrecovery | 0.6 |  | recovery rate to calculate air needed for raw product out of column
Airtotal | 158.7302 | l/min | airflow into bed during adsorption cycle
Aircycle | 36.0911 | l | Volume of air delivered to the bed per cycle
Prodcycle | 4.5475 | l | product delivered per cycle
Purgecycle | 3.0317 | l | product returned to purge per cycle |  |  |  | 30lpm max
Wasteload | 125.3968 | l/min | Average waste collection rate during adsorption
Loadtime | 13.6425 | sec | Time to load the bed with waste
avgblowdown | 125.3968 | l/min | adsorbed waste over bed loading time assuming a/b operation with 2 columns
not adsorbed/cycle | 7.5791 | l/min
recovered | 4.5475 | l/min
purge | 3.0317 | l/min
Raw production from column | 33.3333 | l/min

## Model-tester
MonoColumn VPSA, same bed and pump as O24U
Bed-product recovery ratio
N2 | 0.78 | 0.78 | 0.78 |  |  | Waste | 0.79 | .78 N2 + .01 O2
O2 | 0.209 | 0.209 | 0.209 | 95% |  | Product | 0.21 | 95% O2, 5% Ar
Ar | 0.009 | 0.009 | 0.009 | 4.091%
Other | 0.002 | 0.002 | 0.002 | 0.9091% |  |  | 19.86% | recovered in product
| 1 | 1 | 1 | 0.22 |  |  | 1.145% | adsorbed in bed
|  |  |  | 4.5455 | air/O2
N2cap | 22 | 22 | 22 | l/kg | Li-X ~= 22l/kg
iso-1 | 0.7 | 0.7 | 0.7 |  | fraction of capacity used from P/P0 range read off iostherm chart for adsorb/desorb pressure ratio
cyclecap | 15.4 | 15.4 | 15.4 | l/kg | liters of waste per kg of adsorbent, adsorbed each cycle, no hysteresis
bedweight | 1.8 | 0.5 | 0.5 | kg
bedcap | 27.72 | 7.7 | 7.7 | l | liters of waste adsorbed in the bed each cycle
Prodrate | 33.33 | 20 | 10 | l/min
Prodrecovery | 1 | 1 | 1 |  | recovery rate to calculate air needed for raw product out of column
Airtotal | 158.7143 | 95.2381 | 47.619 | l/min | airflow into bed during adsorption cycle
Aircycle | 35.0886 | 9.7468 | 9.7468 | l | Volume of air delivered to the bed per cycle
Prodcycle | 7.3686 | 2.0468 | 2.0468 | l | product delivered per cycle
Purgecycle | 0 | 0 | 0 | l | product returned to purge per cycle |  |  |  | 30lpm max
Wasteload | 125.3843 | 75.2381 | 37.619 | l/min | Average waste collection rate during adsorption
Loadtime | 13.2648 | 6.1405 | 12.281 | sec | Time to load the bed with waste
avgblowdown | 125.3843 | 75.2381 | 37.619 | l/min | adsorbed waste over bed loading time assuming a/b operation with 2 columns
| 7.3686 | 2.0468 | 2.0468 | not adsorbed/cycle
| 7.3686 | 2.0468 | 2.0468 | recovered
| 0 | 0 | 0 | purge

## model-puffer
MonoColumn VPSA, same bed and pump as O24U
Bed-product recovery ratio
N2 | 0.78 | 0.78 | 0.78 | 0.78 |  |  | Waste | 0.79 | .78 N2 + .01 O2
O2 | 0.209 | 0.209 | 0.209 | 0.209 | 95% |  | Product | 0.21 | 95% O2, 5% Ar
Ar | 0.009 | 0.009 | 0.009 | 0.009 | 4.091%
Other | 0.002 | 0.002 | 0.002 | 0.002 | 0.9091% |  |  | 19.86% | recovered in product
| 1 | 1 | 1 | 1 | 0.22 |  |  | 1.145% | adsorbed in bed
[5 empty] | 4.5455 | air/O2
N2cap | 22 | 22 | 22 | 22 | l/kg | Li-X ~= 22l/kg
iso-1 | 0.7 | 0.7 | 0.7 | 0.7 |  | fraction of capacity used from P/P0 range read off iostherm chart for adsorb/desorb pressure ratio
cyclecap | 15.4 | 15.4 | 15.4 | 15.4 | l/kg | liters of waste per kg of adsorbent, adsorbed each cycle, no hysteresis
bedweight | 1.8 | 0.2 | 0.1 | 0.15 | kg
bedcap | 27.72 | 3.08 | 1.54 | 2.31 | l | liters of waste adsorbed in the bed each cycle
Prodrate | 33.33 | 10 | 5 | 7 | l/min
Prodrecovery | 1 | 1 | 1 | 1 |  | recovery rate to calculate air needed for raw product out of column
Airtotal | 158.7143 | 47.619 | 23.8095 | 33.3333 | l/min | airflow into bed during adsorption cycle
Aircycle | 35.0886 | 3.8987 | 1.9494 | 2.9241 | l | Volume of air delivered to the bed per cycle
Prodcycle | 7.3686 | 0.81873 | 0.40937 | 0.61405 | l | product delivered per cycle
Purgecycle | 0 | 0 | 0 | 0 | l | product returned to purge per cycle |  |  |  | 30lpm max
Wasteload | 125.3843 | 37.619 | 18.8095 | 26.3333 | l/min | Average waste collection rate during adsorption
Loadtime | 13.2648 | 4.9124 | 4.9124 | 5.2633 | sec | Time to load the bed with waste
avgblowdown | 125.3843 | 37.619 | 18.8095 | 26.3333 | l/min | adsorbed waste over bed loading time assuming a/b operation with 2 columns
| 7.3686 | 0.81873 | 0.40937 | 0.61405 | not adsorbed/cycle
| 7.3686 | 0.81873 | 0.40937 | 0.61405 | recovered
| 0 | 0 | 0 | 0 | purge

## Model-VPSA
Ref. MonoColumn VPSA, same bed and pump as O24U
Bed-product recovery ratio
N2 | 0.78 | 0.78 | 0.78 |  |  | Waste | 0.79 | .78 N2 + .01 O2
O2 | 0.209 | 0.209 | 0.209 | 95% |  | Product | 0.21 | 95% O2, 5% Ar
Ar | 0.009 | 0.009 | 0.009 | 4.091%
Other | 0.002 | 0.002 | 0.002 | 0.9091% |  |  | 19.86% | recovered in product
| 1 | 1 | 1 | 0.22 |  |  | 1.145% | adsorbed in bed
|  |  |  | 4.5455 | air/O2
N2cap | 22 | 22 | 22 | l/kg | Li-X ~= 22l/kg
iso-1 | 0.7 | 0.7 | 0.7 |  | fraction of capacity used from P/P0 range read off iostherm chart for adsorb/desorb pressure ratio
cyclecap | 15.4 | 15.4 | 15.4 | l/kg | liters of waste per kg of adsorbent, adsorbed each cycle, no hysteresis
bedweight | 1.8 | 10 | 10 | kg
bedcap | 27.72 | 154 | 154 | l | liters of waste adsorbed in the bed each cycle
Prodrate | 33.33 | 100 | 50 | l/min
Prodrecovery | 1 | 1 | 1 |  | recovery rate to calculate air needed for raw product out of column
Airtotal | 158.7143 | 476.1905 | 238.0952 | l/min | airflow into bed during adsorption cycle
Aircycle | 35.0886 | 194.9367 | 194.9367 | l | Volume of air delivered to the bed per cycle
Prodcycle | 7.3686 | 40.9367 | 40.9367 | l | product delivered per cycle
Purgecycle | 0 | 0 | 0 | l | product returned to purge per cycle |  |  |  | 30lpm max
Wasteload | 125.3843 | 376.1905 | 188.0952 | l/min | Average waste collection rate during adsorption
Loadtime | 13.2648 | 24.562 | 49.1241 | sec | Time to load the bed with waste
avgblowdown | 125.3843 | 376.1905 | 188.0952 | l/min | adsorbed waste over bed loading time assuming a/b operation with 2 columns
| 7.3686 | 40.9367 | 40.9367 | not adsorbed/cycle
| 7.3686 | 40.9367 | 40.9367 | recovered
| 0 | 0 | 0 | purge

## BOMcost
Zeolite cost
| 1 pump slices |  | 70 | $/kg | early models dual pump share
| 1/4 Scale POC | First product |  |  |  | Table demo | MVP Demo | MVP bigger | Full Scale Demo
[5 empty] | O2 lpm for 1 module | 2.5 | 6-7.5 | 12-15 | 25
| 25 | 100 | lpm produced |  | O2 lpm for 2 modules | 10 | 25-30 | 50-60 | 100
| 360 | 1440 | Air pump lpm displacement |  | Zeolite | 500g | 1-2kg | 3-4kg | 6-10 kg
| 3x48, 6x24 | 8x36 | column |  | Column HW | 2x16 PVC | 2.5x20, 3x24 | 3x36, 4x36 | 4x48, 6x36
| 2.5 | 15 | kg zeolite
[5 empty] | Compressor | 50 | 130 | 260 | 800
| 300 | 500 | Motor + coupling |  | Vacuum pump | 50 | 110 | 130 | 260
| 300 | 500 | PDP |  | Column HW | 50 | 160 | 220 | 500
| 220 | 300 | VFD |  | Zeolite | 30 | 90 | 150 | 450
800 | 250 | 1200 | Dryer
| 400 | 800 | Structural (Skid + rack ) |  |  | 180 | 490 | 760 | 2010
| 500 | 1000 | Column HW (Spools, endcaps, clamps, filter plates) |  | BOS | 200 | 500 | 800 | 1000
| 200 | 400 | O2 tank |  |  | 2 | 2 | 2 | 2
| 200 | 150 | Solenoid Valves |  |  | 760 | 1980 | 3120 | 6020
| 200 | 500 | Backpressure regulator
| 500 | 500 | Sensors + Controller
| 175 | 1050 | Zeolite |  | Compressor FA slpm | 40 | 100 | 200 | 350
| 200 | 800 | wiring, air+ vacuum line |  | Vacuum pump slpm | 40 | 48 | 100 | 200
| 3445 | 7700 | build cost
Compressor | 820 | 1300
Dryer | 250 | 1200
Column + valves | 1275 | 3100
Skid | 1100 | 2100
| 3445 | 7700
O24U
|  | 270 | Pump
|  | 200 | Coalescer + post FIlter
|  | 100 | Cooler
|  | 800 | TSA Dryer (Columns, Valves, Heater)
|  | 500 | Structural (Rack + Base + Panels )
|  | 500 | Column HW (Spools, endcaps, clamps, filter plates)
|  | 120 | O2 tank
|  | 200 | PSA Solenoid Valves
|  | 120 | Backpressure regulator
|  | 300 | Gauges and flowmeters
|  | 800 | Sensors + Controller
|  | 250 | Zeolite
|  | 800 | fittings, tubing, wiring
|  | 4960 | build cost

## intercooler
flowrate | 200 | sl/m
pressure | 2.5 | bar
radius | 0.55 | cm
length | 400 | cm
flow volume | 80 | l/m
area | 0.95033 | cm^2
velocity | 1403.0188 | cm/s
tres | 0.2851 | sec
inner surface | 1382.3008 | cm^2
rho | 2.911 | kg/m^3
vd | 1.983e-05 | Ns/m^2
vk | 6.8121e-06 | mm^2/s | cSt
l | 0.0055 | m
R | 5.2566e-07

## Heater
Watts = SCFM x ∆T / 3
SCFM = Standard Cubic Feet Per Minute
∆T = Temperature Rise in Degrees F from the Inlet to the Outlet
28.317
lpm | 50 | 85
scfm | 1.7657 | 3.0017
deltaT | 200 | 200
watts | 117.7149 | 200.1154
0.66929

## Valves
http://www.fnwvalve.com/FNWValve/assets/images/PDFs/FNW/tech_AboutCv.pdf
| Valve Cv | Fill valve | Purge valve | Purge valve | O2 Valve | O2 Delivery | Check Valve | Blowdown | Air Recycle control valve | Backpressure valve | Air Supply Regulator | 3/5 fill | 3/5 blowdown | 3/5 vac | Cv
| Cv | 0.6 | 4.8 | 4.8 | 0.23 | 0.1 | 0.2 | 0.6 | 0.18 | 0.25 | 0.1 | 0.89 | 0.89 | 0.89 | 0.89 | 0.25
SpGr | G | 1 | 1 | 1 | 1.1 | 1.1 | 1.1 | 1.1 | 1 | 1 | 1 | 1 | 1 | 1 | 1.68 | 0.375
DegF | T | 70 | 70 | 70 | 70 | 70 | 70 | 70 | 70 | 70 | 70 | 70 | 70 | 70 | 2.79 | 0.5
PSIG | Pu | 25 | 25 | 2 | 28 | 10 | 28 | 20 | 35 | 25 | 100 | 20 | 19 | 15
PSIG | Pd | 23 | 2 | 0 | 27.7 | 7 | 27 | 5 | 0 | 0 | 28 | 19 | 2 | 13.5
SCFH | Qs | 311.942 | 7224.0683 | 1589.4915 | 46.3014 | 47.0068 | 73.2052 | 682.8686 | 357.0973 | 385.2535 | 444.842 | 307.5779 | 1088.5963 | 346.5868
SCFH | Qc | 844.2945 | 6754.3563 | 2841.2532 | 331.9033 | 83.4743 | 288.6116 | 703.6175 | 317.089 | 351.7894 | 406.5516 | 1094.641 | 1063.0951 | 936.9117
l/min | Q | 147.22 | 3187.696 | 750.1552 | 21.8518 | 22.1847 | 34.549 | 322.2775 | 149.6491 | 166.0258 | 191.8707 | 145.1604 | 501.7242 | 163.5705
SCFM | Q | 5.199 | 112.5726 | 26.4915 | 0.77169 | 0.78345 | 1.2201 | 11.3811 | 5.2848 | 5.8632 | 6.7759 | 5.1263 | 17.7183 | 5.7764
O2 Backpressure regulator
| Valve Cv | Sentry | Tesco -501 | Conoflow | Generant | Fairchild 66BP
| Cv | 0.2 | 0.3 | 0.25 | 0.2
SpGr | G | 1.1 | 1.1 | 1.1 | 1.1
DegF | T | 70 | 70 | 70 | 70
PSIG | Pu | 25 | 25 | 25 | 25
PSIG | Pd | 10 | 10 | 10 | 10
SCFH | Qs | 247.662 | 371.493 | 309.5775 | 247.662
SCFH | Qc | 268.3344 | 402.5016 | 335.418 | 268.3344
l/min | Q | 116.8832 | 175.3249 | 146.1041 | 116.8832
SCFM | Q | 4.1277 | 6.1915 | 5.1596 | 4.1277

## Materials
Viton
Silicone
EPDM
HNBR
PTFE | Check PTFE compatibility

## Things I have built
Date | What | Invent | Design | Build | # | Skills I learned | Notes
1974 | Guitar Amplifier |  | X | X | 1 | Electronic circuit analysis, mechanical construction, electronic construction, design for specific outcome
1975 | Analog analysis software | X | X | X | 1 | Circuit analysis of 3-terminal analog networks, logical reduction, problem abstraction, software design, use of programming languages
1978-1983 | Sport Auto |  | X | X | 2 | Strength of Materials, welding and brazing, mechanical drawing, use of machine tools
1981 | Anti-Surge Controller |  |  | X | 1 | Use of machine tools, layout and planning, service lifetime
1982 | Safety Control System |  | X | X | 1 | PLCs, System design, fault tolerant design, fault modeling, service life optimization, industrial control systems
1983 | RTU Software |  | X | X | 1 | Communication software design, high reliability design
1983 | Compressed Gas meter | X | X | X | 1 | Physics and thermodynamics, confidence bounds estimation, mechanical design, product and BOM design, project cost and budget
1986 | Microprogram Controller | X | X | X | 1 | Computer logic design, PLA design, rigorous timing analysis, design for test, design for manufacture, failure analysis, project planning and schedule commitment
1989 | Fault tolerant computer |  | X |  | 1 | Computer system architecture, high level tradeoffs, full end to end project responsibility
1987-2008 | Celtic harp |  | X | X | 5 | Fine woodworking, physics and acoustics, design with wood, wood finishing techniques
1999 | Mandol |  | X | X | 1 | Fretted instrument design, lutherie craft
1999-2000 | Information Appliance | X | X | X | 2 | Web connected appliance design, battery device design, browser based interaction, embedded Linux, low power constrained hardware
2001 | Internet-connected Photovoltaic generation |  | X | X | 1 | Embedded hardware and Linux, hardware I/O and SCADA integration, serial protocols, processing of measurement data, real-time web server
2003 | Electrolyre | X | X | X | 1 | Patent and trademark processes, optimization of many conflicting variables, problem solving in design
2005 | Internet Media Devices |  | X | X | 1 | Search and indexing, embedded Linux with streaming media, appliance design and demo, network client + server device integration
2008-2012 | Vacuum tube amplifier |  | X | X | 5 | Deep understanding of vacuum tube physics, development of experimental process and inquiry, instrument construction in metal
2010 | Studio Microphone Preamplifier | X | X | X | 1 | Design to performance goals, low noise design and testing, product BOM, PCB design, industrial design
2018 | Climate instrument |  | X | X | 3 | Maker/DIY process, Arduino software, rapid deployment, iterative design
2021 | Gas Separator |  | X | X | 1 | Physics and thermodynamics, experimental process, industrial controls integration, cottage workshop
[6 empty] | What I Learned | Notes
1970-1975 | Audio Amp |  |  |  |  | Basic electronic,s analog, audio and TV circuts, color TV repair, audio amplifier and radio work | Tube amp hacks and builds in HS
1975 | Fortran |  |  |  |  | Numerical techniques for graph analysis and network reduction | NKANAL and PC GUI
1977-1980 | Electrical Systems |  |  |  |  | Residential, commercial, and indistrial electrical systems, low voltage (480V), technical work, installation, and maintenance. Generators, motor controls, and switchgear to ~1000 KW, offshore | Platform Electrician, industrial control systems, mechanical systems, refrigeration and air handling, water pumping,
1978-1980 | Auto Engr |  |  |  |  | Strength of Materials, Automotive engineering, chassis design, millwright techniques, welding and brazing, working with fiberglass | Auto Projects
1980 | Audio design |  |  |  |  | PC board design and fabrication, audio circuits | Switching system for audio demo
1980-1983 | PC Software |  |  |  |  | Software development for PC systems, | PC software development for invoices and gas custody transfer, tank report proxy
1980-1985 | Process Control |  |  |  |  | Process controls and instrumentation, safety systems, rigorous testing methods, field experience, product design, PLC and digital systems, large scale systems
1985-1987 | Computer Arch |  |  |  |  | CPU Microarchitecture, high speed digital circuit design, design for test and reliability, board level produt development
1987-1989 | [5 empty] | Fault tolerant computing, system architecture, quantitative architecture analysis, memory system design (MMU to VSM), system level product development,
1989-1995 | [5 empty] | System architecture across product lines, system modeling and simulation, performance and benchmarking, managing a department
1996-2000 | [5 empty] | Analytic performance analysis an modeling, workload and memory system interaction, hardware and software scalability
2000-2006 | [5 empty] | Computer architecture research, technology and architecture, protocol design and analysis, high performance shared memory system design
2000-2012 | [5 empty] | Audio recording and production, vintage audio gear, recording consoles
2006-2012 | [5 empty] | Advanced server memory systems, solid state storage and scalable servers
2005-2012 | [5 empty] | Professional audio design and analysis, vacuum tube and hybrid, signal quality and noise analysis, microphone amplifiers and signal compressors, design and production
2012-2020 | [5 empty] | Internet of Things, embedded software, APIs and communication protocols, standardization, semantic interoperability

## vacuum +velocity
1/8" sample tube |  |  | 1/4" | 3/8" | 1/2"
0.0625 | in dia |  | 0.35 | 0.44 | 0.56 | in dia
0.79375 | mm radius |  | 4.445 | 5.588 | 7.112 | mm radius
1.9793 | mm^2 |  | 62.0717 | 98.0986 | 158.9035 | mm^2
505.2225 | mm per cc |  | 16.1104 | 10.1938 | 6.2931 | mm per cc
|  |  | 200 | 200 | 200 | l/min
100 | cc/min |  | 20000 | 20000 | 20000 | cc/min
84.2037 | cm/sec |  | 537.0137 | 339.7943 | 209.771 | cm/sec
0.84204 | m/sec |  | 5.3701 | 3.3979 | 2.0977 | m/sec
29.9213 | in atm | 29.9213
29.67 | in vac
760 | mm atm |  | SA-100 pump
6.382 | mm vac (torr) |  | 6.2 | mm
1000 | mbar atm
8.3974 | mbar vac |  | 8.1579 | mbar vac
| 14.7 | atm
| 0.12 | psia
| 8.1633 | mbar
| 6.2041 | mmhg
| 41 | F
| 5 | C

## Column HW
C
| spool | 3x24 | 2 | 49.99 | 99.98 | 2 | 49.99 | 99.98 | 49.99
| cap top | 1/2 NPT | 2 | 19.99 | 39.98 | 2 | 16.79 | 33.58 | 16.79
| capbotton | 1/4 NPT | 2 | 19.99 | 39.98 | 2 | 19.99 | 39.98 | 19.99
| filter plate | 3" | 4 | 24.99 | 99.96 | 4 | 22.99 | 91.96 | 22.99
| clamp | 3" | 8 | 11.99 | 95.92 | 8 | 11.99 | 95.92 | 12.99
| gasket |  | 8 | 3.72 | 29.76 | 8 | 3.2487 | 25.99 | 3.2487
| mount | spool clamp | 4 | 12.99 | 51.96 | 4 | 12.99 | 51.96 | 12.99
[5 empty] | 457.54 |  |  | 439.37

## electrical
I/P Transducer
0.02 | amps |  | 3.5
189 | ohms |  | 25.4
3.78 | volts |  | 88.9
Volts
11 |  |  |  |  | Measured
| Quantity | Ohms | Amps | Watts | Amps | Watts
O2 Out/Press | 1 | 10.6 | 1.0377 | 11.4151 | 1 | 11
Drier Purge out | 1 | 8.5 | 1.2941 | 14.2353 | 1.24 | 13.64
PSA Air In | 1 | 7.1 | 1.5493 | 17.0423 | 1.4 | 15.4
PSA Blowdown | 1 | 7.1 | 1.5493 | 17.0423 | 1.4 | 15.4
BD Diverter | 1 | 7.3 | 1.5068 | 16.5753 | 1.4 | 15.4
Fan | 1 | C | 0.42 | 4.62 | 0.42 | 4.62
MOV | 1 | 30 | 0.36667 | 4.0333 | 0.4 | 4.4
Total |  |  | 7.724 | 84.9636 | 7.26 | 79.86
base
64
[7 empty] | Digits | Codes
[7 empty] | 1 | 64
[7 empty] | 2 | 4096
[7 empty] | 3 | 262144
[7 empty] | 4 | 16777216
[7 empty] | 5 | 1073741824
[7 empty] | 6 | 68719476736
[7 empty] | 7 | 4398046511104
[7 empty] | 8 | 281474976710656

## bom
|  | Description | Manufacturer/Vendor | Manuf. Part ID | Manuf. Description | Consumable
Frame |  | Relay rack 28U | LeGrand
|  | RackBase | Tap Plastic |  | King Star Board 1"
| 4 | Caster |  |  | 3", 300lb
| 4 | 8u panel |  |  | 14x20 custom
| 2 | 5u panel |  |  | 8-3/4x20 custom
| 4 | Base bolts + lock washers + washers
| 4 | Base thread inserts
| 16 | Caster mounting screws
| 32 | Panel mounting screws
Pump |  | Pump
|  | Air inlet filter | Solberg |  |  | Solberg HE14
Filter mounting adapter
Pump inlet hose fitting
Pump outlet tee
pressure relief valve
pump outlet hose fitting
Pump inlet hose fitting
| 4 | Pump inlet hose clamps
Pump inlet hose tee
| 2 | Pump inlet hose sections
| 4 | Pump mounting tabs
| 4 | Pump mounting bolts and nuts
Condenser |  | Condenser core
Condenser frame assy
Condenser fan plate
Condenser fan
| 2 | Condenser mounting brackets
Condenser inlet hose assy
Condenser outlet hose assy
Condenser grille screen
| 4 | Condenser bracket screws
Sep+Filter |  | Separator inlet tee
Separator inlet hose fitting
Separator inlet temperature probe
Separator outlet nipple
Pressure tap tee
Micron filter inlet nipple
Micron filter outlet fitting
Separator inlet mount bracket
Micron filter mount bracket
| 6 | Bracket mounting screws
Backpressure | 2 | Control valve adapter
Control valve inlet nipple
Control Valve
Bacpressure regulator inlet nipple
Bacpkpressure regulator
Backpressure regulator outlet hose adapter
Backpressure outlet hose
| 2 | Backpressure outlet hose clamps
I/P transducer
I/P inlet fitting
I/P inlet pipe
I/P outlet fitting
I/P outlet pipe
Control valve air inlet fitting
Flow Manifold |  | Manifold block
Inlet fitting
Inlet pipe
Mass Flow sensor
instrument air fitting
Pressure sensor
| 4 | Mounting screws for manifold
| 4 | Mounting screws for MFS
Outlet fitting
Outlet pipe
Crossover fittings
Crossover pipe
Column inlet |  | Check valve inlet fitting
Check valve inlet fitting
Check valve outlet fitting
Check valve outlet pipe
Column 2 inlet Tee
Cross column manifold pipe
Column 1 inlet Elbow
Supply and Exhaust | 4 | Solenoid valves
| 4 | Tube union fittings
| 6 | Tube Union Couplers
| 2 | Exhause muffler/bleeders
| 1 | short solenoid bracket
| 2 | long solenoid brackets
| 6 | Solenoid bracket mounting screws
| 8 | Solenoid mounting screws
Columns | 2 | Filter housings
| 4 | Supply and Exhaust elbows
| 2 | Column top hangers
| 4 | Column hanger mounting screws
| 8 | Column mounting screws
| 2 | Column bottom brackets
| 2 | Internal outlet diffusers
| 2 | Internal inlet diffusers
| 2 | Internal media separators
| 5 | Zeolite 13X media kg
| 1 | Activated Alumina media kg
Column O2 outlet | 2 | Column outlet fittings
| 2 | Column outlet pipes
| 2 | Outlet pipe tee adapter
| 2 | Outlet distributor tees
| 2 | Outlet solenoid valves
Outlet solenoid 1 elbow
Outlet solenoud 2 tee
Outlet cross pipe
Outlet flare union coupler
Outlet flare union fitting to purge inlet Tee
| 2 | Solenoid mounting strips and screws
Purge gas | 2 | Purge check valves
Column 1 Elbow
Column 2 Tee
Purge cross pipe
Purge feed pipe 90
Purge flowmeter
Purge Flowmeter inlet street tee
Purge flowmeter outlet fitting Elbow
O2 tank |  | Tube 90 fitting for feed pipe from purge tee
O2 tank feed pipe
O2 tank inlet tee
O2 tank inlet fitting
O2 tank sample tap fitting
O2 tank sample tap needle valve
O2 sample tap hose barb
O2 tank trheaded reducers
O2 tank outlet fitting
O2 tank pressure sensor
O2 tank lower bracket
O2 delivery circuit |  | O2 tank outlet pipe
O2 regulator inlet elbow
O2 regulator 10-32 adapters
O2 regulator
O2 regulator outlet tee
O2 regulator pressure sensor
O2 regulator outler hose barb
O2 flowmater inlet hose
O2 flowmeter
O2 flowmeter outlet hose sections
O2 check valve
O2 outlet filter
O2 delivery adapter
O2 delivery hose clamps
O2 flowmeter mounting bracket
O2 analyzer |  | O2 sensor electrochemical 0-100%
0-1 lpm flowmeter
custom modification to flowmeter
O2 analyzer mounting bracket
Controls |  | IEC fused filtered power inlet
IEC power cable US
18/3 SO cable for power and motor feed
DIN rail
DIN rail terminal block
DIN rail power supply 5V
DIN rail power supply 12V
DIN rail primary AC relay
DIN rail raspberry Pi enclosure
DIN rail RTU assembly
| 6 | DIN rail DC SSR for solenoid valve
DIN rail DC SSR for fan control
DIN rail AC SSR for pump control
DIN rail PID controller multi-input
Finger raceway
Mounting screws for raceway and rail
Hookup wire
USB cable for RTU
USB power cable for Rpi
RS-485 USB stick
Raspberry Pi 4
User controls |  | Control box
USB cable for control box
Arduino board
2x20 VFD
3 position switch with N.O and N.C blocks
| 3 | status indicator light green/yellow/red
Internal USB cable type B bulkhead to micro
RTU |  | RTU internal
Arduino pro mini or ESP32
Instrumentation amplifier
USB cable type B to micro internal
Micro RTU |  | Arduino Pro Mini
USB stick case
5 pin connector
USB cable
Misc |  | cable clamps and wire handling
wire markers
extension connections for solenoids

## EquipmentList
Equipment | PWM RTU
Point | AC RTU
Air RTU
O2 RTU
Air side | ID | P&ID Tag | Description | BOM
Air Pump | AP101 | AP-101 | Air Pump
Air Pump | M101 | M-101 | Motor for air pump
Inet filter | X101
Relief valve | V101
Condenser | E101
Condenser Fan | M102 | M-102
Separator inlet temp | T101 | TE-101
Separator | D101
Column 1 | C110 |  | H2O Adsober 1
Inlet valve | V110 | V-110
Outlet check | V111
Regen in check | V112
Regen out valve | V113 | V-113 | AC Solenoid
Regen out temp | T110 | TE-110
Column 2 | C120 |  | H2O Adsorber 2
Inlet valve | V120 | V-210
Outlet check | V121
Regen in check | V122
Regen out valve | V123 | V-123 | AC Solenoid
Regen out temp | T120 | TE-120
Regen Air flow | F101 | FIC-101
Regen N2 Flow | F102 | FIC-102
Regen Heater | H101 | IC-101 | PWM via internal RTU
Regen in temp | T102 | TE-102
BP Regulator | V102
Back Pressure Valve | V103 |  | 3-6 (?) PSI control range
I/P transducer | X103 | XV-103 | 4-20 mA to 3-15 psig
I/P Control Pressure | P102 | PI-102
Micron FIlter | X102
Air outlet flow | F103 | FIC-103
AIr Outlet Pressure | P101 | PI-101
O2 Side
Air inlet flow | F201 | FT-201
Air inlet temp | T201 | TE-201
Air Inlet Pressure | P201 | PT-201
Column 1 | C220 |  | N2 Adsober 1
Air inlet valve | V220 | V-220 | AC Solenoid
N2 Purge Valve | V221 | V-221 | AC Solenoid
O2 out check | V222
O2 purge check | V223
Column 2 | C230 |  | N2 Adsober 2
Air inlet valve | V230 | V-230 | AC Solenoid
N2 Purge Valve | V231 | V-231 | AC Solenoid
O2 out check | V232
O2 purge check | V233
O2 purge valve | V201 | V-201 | AC Solenoid
O2 Purge Flow | F202 | FIC-202
O2 repressure valve | V202 | V-202 | AC Solenoid
O2 Accumulator | D201
O2 Accum pressure | P202 | PI-202
O2 Accum pressure | P202 | PT-202
O2 Regulator | V203
O2 Delivery Pressure | P203 | PI-203
O2 Delivery Pressure | P203 | PT-203
O2 Delivery Flow | F203 | FT-203
O2 Delivery Flow | F203 | FIC-203
O2 Delivery Quality | Z201 | Z-201 | %O2 Sensor
Blowdown Flow | F204 | FIC-204
O2 Column Sample Flow | F205 | FIC-205
O2 Accum Sample Flow | F206 | FIC-206
O2 Delivery | F207 | FIC-207
O2 Production | F208 | FIC-208
Blowdown regulator | V203 | BPV-203
Blowdown Pressure | P204 | PI-204
Blowdown Bypass | V204 | V-204 | AC Solenoid

## RTU config
AC RTU | [19 empty] | AC RTU | Ref | Func | Ribbon Pin | MCU Pin | I2C address
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 |  |  |  |  | C220 Inlet | V220 | Open | 2 | D32
Top row | L | N | L | N | L | N | L | N | L | N | L | N | L | N | L |  |  |  |  | C220 blowdown | V221 | Open | 4 | D34
| V220 |  | V221 |  | V230 |  | V231 |  | V201 |  | V202 |  | SP |  | OUT |  |  |  |  | C230 Inlet | V230 | Open | 6 | D36
[20 empty] | C230 blowdown | V231 | Open | 8 | D38
Bottom row | V113 |  | V123 |  | V204 |  |  |  | NC |  |  | STBY |  | AC | IN |  |  |  |  | O2 Purge | V201 | Open | 10 | D40
| L | N | L | N | L | N | [5 empty] | - | + | N | L |  |  |  |  | O2 Repressure | V202 | Open | 12 | D42
[20 empty] | C110 Purge Out | V113 | Open | 1 | D33
Air RTU | [19 empty] | C120 Purge Out | V123 | Open | 3 | D35
| 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 |  | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18 | Blowdown Backpressure | V204 | Close | 5 | D37
Top row | + | - |  |  |  | - | - | + | + |  | + | - | NC | + | - | NC | + | - | NC | Spare |  |  | 14 | D44
| V103 |  |  | NC |  |  |  | T101 |  |  |  | T102 |  |  | T110 |  |  | T120
[20 empty] | Local display |  | scl |  | D21 | 0x78
Bottom row | M101 |  | M102 |  |  | V110 | [5 empty] | V120 |  |  |  | NC |  | DC |  |  |  | sda |  | D20
| + | - | + | - | + | - | OP | CL | COM |  | + | - | OP | CL | COM |  |  | + | -
Air Control RTU
[20 empty] | Pump Ctl | M101 | Run | 5 | D11
[20 empty] | Fan Pwr | M102 | Fwd | 7 | D10
[22 empty] | Rev | 9 | D9
[20 empty] | C110 Inlet | V110 | Vopen | 11 | D8
[22 empty] | Vclose | 13 | D7
[22 empty] | Sopen | 19 | D26
[22 empty] | Sclose | 21 | D25
[20 empty] | C120 Inlet | V120 | Vopen | 15 | D6
[22 empty] | Vclose | 17 | D5
[22 empty] | Sopen | 23 | D24
[22 empty] | Sclose | 25 | D23
[20 empty] | Air Bypass Control | V103 | PWM | 6 | D4
[20 empty] | Cond Temp | T101 | CLK | 8 | SCK
[22 empty] | SDO | 10 | MISO
[22 empty] | SDI | 12 | MOSI
[22 empty] | CS | 14 | D53
[22 empty] | RDY | 16 | D52
[20 empty] | (SPI) |  | SDO | 18 | D27
[20 empty] | (SPI) |  | CLK | 20 | D28
[20 empty] | Regen Temp | T102 | CS | 22 | D29
[20 empty] | C110 temp | T110 | CS | 24 | D30
[20 empty] | C120 temp | T120 | CS | 26 | D31
[20 empty] | Local display |  | scl |  | D21 | 0x78
[22 empty] | sda |  | D20

## O2 RTU
O2 RTU
|  | Interface | Pins | Conn pins | Func | MCU | I2C address
Air inlet flow | F201 | I2C | 4 | 4 |  |  | 0x40
[5 empty] | RES/L | D22
O2 Delivery Flow | F203 | I2C | 4 | 4 |  |  | 0x02
[5 empty] | SCL | 21
[5 empty] | SDA | 20
Air inlet temp | T201 | RTD | 3 | 3
[5 empty] | CLK | SCK
[5 empty] | SDO | MISO
[5 empty] | SDI | MOSI
[5 empty] | CS | D53
[5 empty] | RDY | D52
Air Inlet Pressure | P201 | 1-5V | 3 | 3 |  | A0
O2 Accum pressure | P202 | 1-5V | 3 | 3 |  | A1
O2 Delivery Pressure | P203 | 1-5V | 3 | 3 |  | A2
O2 Quality Analyzer | Z201 | AIN | 3 | 3 |  | A3
Run Switch |  | DIN | 2 | 3 |  | D6
O2 Status Indicator |  | DOUT*3 | 4 | 4 | ALM/L | D7
[5 empty] | LOW/L | D8
[5 empty] | OK/L | D9
VFD |  | SPI | 5 | 5
[5 empty] | SIO | D23
[5 empty] | STB/L | D24
[5 empty] | SCK | D25
Standby Switch |  | DC | 2 | 3
Standby Switch out |  | DC | 2 | 3
|  | Total | 38